## Sheet1
Annex 3. Summary of analytical conditions forSr isotope analysis by LA-ICP-MS
Sample Preparation
Laboratory name | Geological Survey of Norway
Sample type/mineral | carbonate
Laser ablation system
Make, Model & type | Teledyne-Cetac Analyte Excite
Ablation cell & volume | HelEx II two volume cell
Laser wavelength | 193 nm
Pulse width | < 4 ns
Fluence | 4 J.cm-2
Repetition rate | 10 Hz
Ablation duration | 50 sec
Spot diameter | 85 µm
Sampling mode / pattern | static spot ablation
Carrier gas | 100% He in the cell, Ar make-up gas combined using a T-piece along the sample transport line. He/Ar in torch typically 0.4/1
Cell carrier gas flow (l/min) | ~0.7 l/min
ICP-MS Instrument
Make, Model & type | Nu instruments Plasma 3 MC-ICP-MS
Sample introduction | ablation aerosol
RF power | 1300 W
Make-up gas flow | ca 0.9 l/min Ar
Detection system | MC on Faraday cups with 1011 Ω resistors
Masses measured | 82, 82.5, 83, 83.5, 84, 84.5, 85, 85.5, 86, 86.5, 87, 87.5, 88
Integration time per peak/dwell times | 0.1 s
Total integration time per output datapoint | 0.1 sec
Data Processing
Gas blank | Each analysis included 35 seconds of gas blank measurement, which is subtracted from each measurement
Calibration strategy | marine mollusk as primary reference material; BCR2 for accurate Rb fractionation correction and marine foramimifera, chalk and basaltic fused glass BHVO-2 as secondary/validation reference materials
Reference Material info | Marine mollusk presumed modern day Sr isotopic composition ( 87Sr/86Sr = 0.70918, McArthur et al. 2001); marine foraminifera reference value of 0.70718 (Thomsen & Andreasen (2019)), chalk reference value of 0.70786, Thomsen & Andreasen (2019)). basaltic glass BHVO-2G reference value of 0.703487 (Weiss et al. 2006) and BCR-2 0.705019 (Weiss et al. 2006).
Data processing / Correction for LIEF | Iolite 4 with DRS "Sr isotopes universal" for data processing including baseline substraction and drift during the analytical session (Mulder et al. 2023)
Mass discrimination | Mass bias for Sr is calculated based on the Kr corrected 86Sr/88Sr for the primary reference material and applied with an exponenctial law, mass bias for Rb is based on the basaltic glass reference material BCR-2G and also applied with an exponential law (see Mulder et al. 2023).
Interference corrections | Kr and Rb interference correction. No corrections were performed for REE++ due to low abundance
Uncertainty level & propagation | reported uncertainties are 2SE internal.
Quality control / Validation | marine foraminifera: 87Sr/86Sr value of 0.709188 ± 0.000017 (MSWD = 0.26, n = 54, reference value of 0.70718, Thomsen & Andreasen (2019)), chalk: 0.707878 ± 0.000028 (MSWD = 0.41, n = 54, reference value of 0.70786, Thomsen & Andreasen (2019)); basaltic glass BHVO-2G: 87Sr/86Sr value of 0.70363 ± 0.00014 (MSWD = 0.16, n = 54, 0.703487 (Weiss et al. 2006))